## Slide 1: Effect of valve holding chamber on the dynamics of aerosol: Time resolved velocimetry study.
Brenda Hervieu1, Anthony Besq2, Jean-Regis Angilella2 & Thierry Porée1
1 OptimHal Protec’Som, Valognes 50700, France
2 Normandie Université, UNICAEN, UNIROUEN, ABTE, Caen, 14000, France

## Slide 2
pMDI is used in the treatment of many respiratory diseases:
  Hand-lung coordination,
  High-speed velocities,
  Oropharyngeal impaction.
VHC is commonly used to improve the efficiency of drug delivery by reducing the need for good hand-lung coordination and limiting the oropharyngeal impact:
  Decrease velocity,
  Retain drug particles.
1

## Slide 3
Success of inhalation therapy depends on several factors, including the patient, and the size and velocity of drug particles.
The aim of this study focuses on the dynamics of aerosols emitted in a VHC according to the inspiratory flow rate, using as a reference the free jet which is the jet emitted without an inhalation chamber.
The dynamics were studied using time-resolved particle image velocimetry (PIV) for two different formulations:
  a solution of beclomethasone dipropionate (BDP),
  a suspension of fluticasone propionate (FP).
2

## Slide 4
3

## Slide 5
4
Simultaneous acquisition
Delay between shaking and actuation

## Slide 6
Paire of images
=> Query windows
Cross correlation between area
1
2
Velocity field
3
1
2
3
PIV is based on the recording of successive images of particles which was considered like a tracers illuminated by a laser beam.
  High Speed camera (Phantom v2640),
  MESA PIV (532-80L) Nd:YAG pulsed laser (wavelength of 532 nm) with a repetition rate up to 30 kHz and a pulse width lower than 150 ns.
5

## Slide 7
6

## Slide 8
: temporal evolution
The temporal evolution of the maximum velocity recorded for the triplicates:
  Five stages
    1: rapid increase, max 78 m/s
    2: slowdown of about 18 m/s
    3: increase, 70 m/s
    4: slow decrease
    5: rapid decrease
Emission duration lasts 130 ms
7

## Slide 9
The temporal evolution of the maximum velocity present five stages.
The temporal scenario remains similar.
Maximum velocities at 15 and 30 L/min are similar – 10 m/s lower than the free jet.
8
: temporal evolution

## Slide 10
The temporal evolution of the maximum velocity recorded for the triplicates:
  Four stages
    1: rapid increase, max 34 m/s
    2: slowdown of about 19 m/s
    3: increase, 30 m/s
    (longer than FP Δ= 80 ms)
    4: decrease
Emission duration lasts 170 ms
9
: temporal evolution

## Slide 11
The temporal evolution of the maximum velocity presents four stages.
No difference for the maximum velocities for the free jet or at 15 and 30 L/min.
10
: temporal evolution

## Slide 12
 | FP | BDP
Maximum Velocity | 78 m/s _ 70 m/s | 34 m/s _ 30 m/s
Flow rate and chamber influence | Decrease | No influence
Jet duration | 130 ms | 170 ms
Stages duration | Shorter | Longer
11
: temporal evolution

## Slide 13
Time average velocities (20 ms)
12
: Velocity magnitude fields
Free jet:
Axisymmetric,
Mean velocities 50 m/s,
Decrease along both x and y axes,
At 3 cm = 30 m/s.

## Slide 14
Free jet:
Axisymmetric,
Mean velocities 50 m/s,
Decrease along both x and y axes,
At 3 cm = 30 m/s.
In a chamber:
Mean velocity decrease 35 m/s,
Decrease along both x and y axes,
At 3 cm = 20 m/s,
Shifted toward the lower part.
13
Time average velocities (20 ms)
: Velocity magnitude fields

## Slide 15
14
Time average velocities (20 ms)
: Velocity magnitude fields
Free jet:
Non symmetric,
Mean velocities 8 m/s (lower part),
Decrease along both x and y axes.

## Slide 16
Free jet:
Non symmetric,
Mean velocities 8 m/s (lower part),
Decrease along both x and y axes.
In a chamber:
Non symmetric,
Mean velocity increase 12 m/s (lower part),
Decrease along both x and y axes.
15
Time average velocities (20 ms)
: Velocity magnitude fields

## Slide 17
Free jet:
Two formulations with the same temporal pattern but different characteristics (nozzle, composition)
  Velocity,
  Duration.
In a chamber:
Not affect the temporal pattern,
Velocity can increase (BDP) or decrease (FP) with the co-flow.
 | FP | BDP
Maximum Velocity | 78 m/s _ 70 m/s | 34 m/s _ 30 m/s
Flow rate and chamber influence | Decrease | No influence
Jet duration | 130 ms | 170 ms
Stages duration | Shorter | Longer
16

## Slide 18
Thank you for your attention